## Sheet1
LPC Checkout
Instrument S/N:0006
Date:6/5/24
By:Mark Turner
Physical Checkout | Y/N | Notes
All parts installed? | Y
Threaded inserts in all plastic parts? | Y
All fasteners installed and hand tight? | Y
Cable harnesses installed? | Y
Cable harnesses laced and neatly routed? | Y
Tubing installed? | Y
Correct tubing for each application? | Y
Electrical Checkout
Verify grounds between power connector, Main board, PHA? | Y
Verify main board revision G, fully checked out? | Y
Verify PHA revision E, Note calibration file? | Y | 1
SD Card installed in both Teensys? | Y
*Power on instrument on bench power supply*
Power supply voltage (~ 16V) | 15 | Volts
Current before Optical head start up | 60 | mA
Current after optical head start up | 190 | mA
Current with pumps running | 1000 | mA
*Flash MB teensy with OPC_V5_1*
Connect MB to laptop running LOPC Python Interface
Pump calibration and flow check
*Connect TSI flow meter to inlet and step through pump settings*
Pump BEMF Set point | Pump1 I | Pump2 I | Pump1 PWM | Pump2 PWM | Instrument I | Instrument Flow | TSI Flow
6.5 | 175.82 | 183.15 | 477 | 475 | 0.61 | 9.19 | 9.51
7 | 219.78 | 234.43 | 538 | 535 | 0.74 | 10.1 | 10.38
7.2 | 249 | 264 | 561 | 556 | 0.78 | 10.33 | 10.73
7.4 | 271 | 285 | 585 | 580 | 0.83 | 10.74 | 11.13
7.6 | 286 | 300 | 606 | 607 | 0.89 | 11.07 | 11.46
7.8 | 330 | 344 | 632 | 630 | 0.95 | 11.43 | 11.8
8 | 366 | 366 | 658 | 655 | 1.02 | 11.62 | 12.12
8.2 | 410 | 425 | 688 | 684 | 1.11 | 12.02 | 12.52
8.5 | 464 | 462 | 724 | 720 | 1.23 | 12.44 | 12.99
9 | 542 | 542 | 782 | 782 | 1.41 | 13.34 | 13.73
Aerosol Calibration
Install aerosol filter on inlet and verify aerosol counts | 1 | #/sample
*Connect PAGS to inlet, step through PSL with total particle counts between 10,000 and 25,0000, select idefault instrument
*For each PSL size (from 300nm - 2000nm) save a 60s data file
PSL Diameter | Approx Peak Diam | Counts/sample | Filename