## Sheet1
LPC Checkout
Instrument S/N: | 7
Date: | 2024-06-05
By: | Mark Turner
Physical Checkout | Y/N | Notes
All parts installed? | Y
Threaded inserts in all plastic parts? | Y
All fasteners installed and hand tight? | Y
Cable harnesses installed? | Y
Cable harnesses laced and neatly routed? | Y
Tubing installed? | Y
Correct tubing for each application? | Y
Electrical Checkout
Verify grounds between power connector, Main board, PHA? | Y
Verify main board revision G, fully checked out? | Y
Verify PHA revision E, Note calibration file? | Y | 5
SD Card installed in both Teensys? | Y
*Power on instrument on bench power supply*
Power supply voltage (~ 16V) |  | Volts
Current before Optical head start up | 70 | mA
Current after optical head start up | 170 | mA
Current with pumps running | 1.16 | mA
*Flash MB teensy with OPC_V5_0*
Connect MB to laptop running LOPC Python Interface
Pump calibration and flow check
*Connect TSI flow meter to inlet and step through pump settings*
Pump BEMF Set point | Pump1 I | Pump2 I | Pump1 PWM | Pump2 PWM | Instrument I | Instrument Flow | TSI Flow
6.5 | 344 | 227 | 517 | 560 | 0.85 | 9.64 | 9.85
7 | 417 | 293 | 569 | 612 | 0.98 | 10.48 | 10.85
7.2 | 432 | 322 | 594 | 64 | 1.06 | 10.93 | 11.25
7.4 | 476 | 351 | 623 | 668 | 1.13 | 11.3 | 11.68
7.6 | 483 | 403 | 649 | 686 | 1.21 | 11.63 | 11.95
7.8 | 535 | 454 | 687 | 713 | 1.31 | 11.84 | 12.25
8 | 578 | 498 | 718 | 733 |  | 12.05 | 12.61
8.2 | 615 | 549 | 742 | 772 | 1.49 | 12.41 | 12.85
8.5 | 674 | 615 | 786 | 805 | 1.64 | 12.97 | 13.35
9 | 769 | 740 | 855 | 866 | 1.85 | 13.52 | 14.07
Aerosol Calibration
Install aerosol filter on inlet and verify aerosol counts | 1 | #/sample
*Connect PAGS to inlet, step through PSL with total particle counts between 10,000 and 25,0000, select idefault instrument
*For each PSL size (from 300nm - 2000nm) save a 60s data file
PSL Diameter | Approx Peak Diam | Counts/sample | Filename
300 | 300 | 22000 | LPC0007_300nm
490 | 400 | 16000 | LPC0007_490nm
890 | 690 | 27000 | LPC0007_890nm
1069 | 980 | 14500 | LPC0007_1069nm