## Sheet1
LPC Checkout
Instrument S/N: | 0008
Date: | 2024-06-06
By: | Mark Turner
Physical Checkout | Y/N | Notes
All parts installed? | Y
Threaded inserts in all plastic parts? | Y
All fasteners installed and hand tight? | Y
Cable harnesses installed? | Y
Cable harnesses laced and neatly routed? | Y
Tubing installed? | Y
Correct tubing for each application? | Y
Electrical Checkout
Verify grounds between power connector, Main board, PHA? | Y
Verify main board revision G, fully checked out? | Y
Verify PHA revision E, Note calibration file? | Y | 1
SD Card installed in both Teensys? | Y
*Power on instrument on bench power supply*
Power supply voltage (~ 16V) | 15 | Volts
Current before Optical head start up | 60 | mA
Current after optical head start up | 190 | mA
Current with pumps running | 810 | mA
*Flash MB teensy with OPC_V5_1*
Connect MB to laptop running LOPC Python Interface
Pump calibration and flow check
*Connect TSI flow meter to inlet and step through pump settings*
Pump BEMF Set point | Pump1 I | Pump2 I | Pump1 PWM | Pump2 PWM | Instrument I | Instrument Flow | TSI Flow
6.5 | 146.5 | 139 | 469 | 467 | 0.58 | 6.9 | 7.36
7 | 190 | 190 | 525 | 520 | 0.68 | 7.57 | 8.08
7.2 | 205 | 205 | 546 | 543 | 0.71 | 7.8 | 8.37
7.4 | 220 | 227 | 568 | 562 | 0.73 | 8.15 | 8.68
7.6 | 234 | 250 | 590 | 584 | 0.81 | 8.32 | 8.89
7.8 | 264 | 278 | 613 | 603 | 0.86 | 8.6 | 9.21
8 | 286 | 300 | 638 | 624 | 0.89 | 8.73 | 9.43
8.2 | 315 | 330 | 662 | 644 | 0.95 | 9.1 | 9.69
8.5 | 344 | 381 | 699 | 683 | 1.04 | 9.41 | 10.09
9 | 418 | 462 | 763 | 740 | 1.22 | 9.89 | 10.71
Aerosol Calibration
Install aerosol filter on inlet and verify aerosol counts | 1 | #/sample
*Connect PAGS to inlet, step through PSL with total particle counts between 10,000 and 25,0000, select idefault instrument
*For each PSL size (from 300nm - 2000nm) save a 60s data file
PSL Diameter | Approx Peak Diam | Counts/sample | Filename
300 | 300 | 16500 | LPC0008_300nm
490 | 490 | 12000 | LPC0008_490nm
890 | 950 | 21500 | LPC0008_890nm
1069 | 1150 | 15000 | LPC0008_1069